## Slide 1: Sample Prep Sorting: Cytek Aurora CS cell sorter

## Slide 2: Cytek Aurora CS cell sorter
The Cytek Aurora CS a sterile cell sorter. It can currently analyze up to 40 colors and sort up to 6 populations at one time. Samples are run by core personnel.
The cell sorter is housed in a BioBubble allowing for the sorting of BSL-2 samples. No BSL-3 samples can be sorted.
Samples do not need to be fixed to run on this instrument.
  However, source of the samples should be disclosed to the flow personnel.

## Slide 3: Aurora CS sorting samples
Recommended sample concentration is 10 million cells/ml and adjust according to your cells count with a minimum final volume of 200 ul.
For final volume, cells need to be diluted in sorting buffer. Recipe can be found here: https://medicine.uams.edu/mbim/research-cores/flow-cytometry-core-facility/protocols-and-reagents/sorting-and-staining-buffers/
If cells are over diluted, it causes unnecessarily long sort times and too concentrated samples causes poor sorting efficiency. You can bring extra sorting buffer with you so that cells can be diluted, if necessary, by core personnel.
Filter cells through blue cap filter tubes (BDFalccon #352235) prior to bringing to the core for sorting. If any clumps are visible, then cells will not be sorted

## Slide 4: Aurora CS sorting samples
Bring collection tubes you want to sort live cells into. We can sort into 96 well plates, 5 ml FACS tubes, 15 ml conical tubes, and 1.5 ml microcentrifuge tubes. If sorting more than 2 populations at once, 15 ml tubes cannot be used. Up to 6 populations can be sorted using 5ml and 1.5ml tubes.
Prepare collection tubes with 0.5-1 ml of media or solution of your choice in the collection tubes to preserve viability. It is recommended for low cell number to also coat the insides of the collection tubes to prevent cells sticking to sides and drying out. Ideally you want to use the collection tubes that best fit the number of cells you expect to get back.
The 1.5ml tubes hold about 200-300k cells, 5 ml FACS tubes hold 1 million or less. If you plan to collect over 1 million sorted cells, then it’s best to use 15 ml conical tubes for collection.
To coat the tubes, fill them with 4% BSA in 1X PBS (filter-sterilized) and incubate at 4C for at least 1 hour prior to sorting or overnight. Just prior to bringing the tubes for sorting, pour out the BSA and then add about 0.5 ml of media to the tubes.
Polypropylene tubes are recommended for collection. The sorter charges the stream to sort cells and a charge may build around polystyrene tubes as the cells are sorted. Initial samples may be in either polystyrene tubes or polypropylene tubes.
After sorting, it is recommended the cells rest on ice for minimum of 30-60 minutes before moving to next step as the sorting procedure puts the cells under strain and stress. Not resting after the sort can lead to increased cell death.

## Slide 5: Preparing Cells
All cell types require different digestion, harvesting, and isolation techniques. It is recommended you find a protocol specific for your cell type that gives you a single cell suspension with optimal cell number recovery.
COUNT YOUR CELLS! This step is very important. After cell collection and prior to staining, count your cells. This will help you determine antibody concentrations, as well as determine if your cell harvest was successful and you have enough cells to stain.
Make sure you have a single cell suspension. If you see clumps in your sample, these cannot be run on the cytometer. They can damage the instrument and cause inaccurate results.
  All samples should be filtered through a 40μm cell strainer topped 5 ml FACs tube before bringing the samples to the flow core.

## Slide 6: Samples
Samples must be in 5ml round bottom FACS tubes or 15 ml conical tubes to be ran on Cytek aurora CS.
The recommended sample concentration is 10 million cells/ml and adjust according to your cell count with a minimum final volume of 200 ul. If cells are over diluted, it causes unnecessary long run times.
Filter cells prior to bringing to the core. If aggregates are present, the samples cannot be ran as it will clog the instrument. They MUST be filtered through blue caps tubes immediately prior to sorting.
Include a sheet with samples indicating all necessary info: Lab name, your name, contact info, staining panel including markers & colors used, cell type, etc. This must be submitted with samples, or it can be emailed to core personnel ahead of time.
Make sure tubes are numbered or labeled legibly to prevent mishandling of tubes out of order.
For sorting, it is recommended samples are brought on ice and covered.

## Slide 7: Panel Design
The Cytek Cloud website offers a variety of tools to build panels for the full spectrum instruments. Visit https://cloud.cytekbio.com/newlogin and create a user account. It will request the instrument serial number to create an account. That is S0318.
Use the Spectrum Viewer to build your panel and determine if the panel is good with a low similarity index.

## Slide 8: Experimental Quality Controls
Quality control samples are needed with each experiment. Best practice is to prepare these with each run. A multi-color experiment cannot be run without proper controls.
For best practice, the controls needed for each experiment are as follows:
  Negative/unstained cells
  Single stain controls for each color including viability dye
  If only staining/sorting for a single color, and no viability dye, then just a negative unstained control is required.
VERY IMPORTANT: The full spectrum cytometer perform an auto fluorescence/background subtraction which yields better results. To properly perform this, we must have unstained cells of each cell type you are analyzing. You cannot substitute other cell types for the unstained sample.
If sorting for one color and adding a viability dye then controls needed are: 1-Unstained cells, 2-cells with viability dye only, 3-cells with pos+ marker only
If you are using the same panel for multiple days of an experiment then settings can be re-used
These controls are needed to set instrument voltages and compensate the samples, not bring these controls can result in poor data.

## Slide 9: Experimental Quality Controls
Some cells types are prohibitive in the number of cells available, making it difficult to generate control samples. In these instances other options are available. Many researchers will use compensation beads to stain single stain controls.
VERY IMPORTANT: For full spectrum flow cytometry, the marker used on the compensation controls, must match exactly the marker used on the samples.
Make sure that the sample/cells chosen for the single stain controls will express your marker. There has to be a positive population for the marker to accurately compensate. If you are unsure that the marker you are using is expressed by the cells in your assay, then use stained compensation beads. Can be found here: https://www.thermofisher.com/order/catalog/product/01-2222-41

## Slide 10: Experimental Quality Controls
The single stain/compensation controls are used to set up the instrument and samples cannot be run without them.
Example: Panel stained with
FITC=CD3, PE=CD4, APC=B220, and EF780=viability dye
Tube 1. Cells only
Tube 2. Cells stained with FITC CD3 only
Tube 3. Cells stained with PE CD4 only
Tube4. Cells stained with APC B220 only
Tube 5. Cells stained with ef780 viability only

## Slide 11: Experimental Quality Controls
FMO controls are used for analysis purposes. When drawing gates to determine which proportion of the cell population expresses a given marker, these controls are used to help set the positive and negative gates. While some populations have clear pos/neg expression, low expressing markers or markers that just cause a slight shift are easier to confidently gate on with FMO controls
FMO example: Using the same panel listed on previous slide FMO controls would be:
Tube 6: FITC FMO- Cells with all stains except FITC
Tube 7: PE FMO- Cells with all stains except PE
Tube 8: APC FMO- Cells with all stains except APC
Tube 9: ef780 viability FMO- Cells with all stains except ef780

## Slide 12: Staining Protocol
All cell types require different variations to the staining protocol, so it’s best to find protocols specific to your assay/cell type.
There are many protocols available on-line and on antibody manufacturer websites.
A viability dye is always needed for sorting cells. Commons ones to use are PI, DAPI and 7AAD.
It is highly recommended that you test your protocols on a small sample size prior to staining a large number of cells/samples on valuable cells. Keep it simple to start.
A generic staining protocol can be found on my website here:
https://medicine.uams.edu/mbim/research-cores/flow-cytometry-core-facility/sample-submission/sample-preparation/

## Slide 13: Analysis and Data
The Aurora uses SpectroFlo software to collect data. These files will be uploaded onto the Flow server for the user to then analyze in FlowJo or similar software. You are able to download FlowJo from their website and the core has a USB dongle you can check out free of charge.
Data Delivery
All data files are uploaded to a folder on a UAMS server under the Flow Core or to a shared folder on UAMS BOX drive.